## Slide 1: MEBT Quadrupoles
Scott Lawrie
FETS Meeting, RAL
17th December 2014

## Slide 2: Yoke and Coil
September
October

## Slide 3: Visual Inspection
Discolouration of a dab of superglue when resin heated in curing process
Some copper badly discoloured, but apparently does not affect voltage holding

## Slide 4: First Completed Assembly
November

## Slide 5: Magnetic Field Mapping
Quadrupole
Dipole
December

## Slide 6: Magnetic Field Mapping

## Slide 7: Field Homogeneity
Magnetic Length = 83.0 mm
Gradient
= 20.6 T/m
Linear to 0.2%

## Slide 8: Field Homogeneity
Homogeneity = 0.2%
Bad horz. hom.  not correctly centred!

## Slide 9: Excitation Linearity

## Slide 10: Harmonic Content
Optimisation Simulations
Measurements of Quad #2

## Slide 11: Summary Statistics
Property | Requirement | Measured
Quadrupole Gradient | 20.0 T/m @ 210 A | 20.6 T/m @ 216 A
Dipole Centre Field | 12 mT @ 8.5 A | 12 mT @ 9.0 A
Integrated Field | 1.70 T | 1.71 T
Magnetic Length | 85 mm | 83 mm
Gradient vs. Current Linearity | 1% saturation at 20 T/m | 3% saturation at 20 T/m
Integrated Field Gradient Homogeneity | 0.5% both planes | 0.2% both planes, (assuming corrected horizontal centre)
Harmonic Content | <0.1% all harmonics | <0.06% all harmonics
Next Steps: - All test documentation collated this week
- Packaged up for Santa’s delivery next week